## Slide 1: Trunk Tracking on NOARK
Stereo ICP vs SAM 3D Body vs a distilled student: why ICP fails in the game, why the student overfit, and why we render synthetic humans

## Slide 2: What we measure
One stereo camera (Radxa, 2 × fisheye) on the table, facing the seated person
Truth: Motive rigid body on the sternum, GPIO-synced to the camera
Takes per person
  woc: instructed trunk sweeps, no compensation
  wc: reaching with trunk compensation
  game: RandomReach with the robot cube
8 people: 5 used for training, 3 never seen (Arjun, Jerome, Lee)
Target: Radxa Dragon Q6A, ≥ 10–15 frames/s, no extra setup
F
Flexion
Lean forward / back
L
Lateral bend
Lean to the side
A
Axial rotation
Twist about the spine
T
Total rotation
Angle of the single rotation from the start pose; needs no camera-to-mocap axis fit
Notes: Every number in this deck is an angle in degrees, compared against Motive mocap of a sternum rigid body. Total rotation needs no axis alignment, so it is the cleanest single number.

## Slide 3: What the takes look like
woc trunk sweeps
wc reaching, compensating
game no compensation
game_wtc with compensation
Arjun
Jerome
Lee
New people · 3× speed · click to play
Notes: Reference camera, rotated upright, 3x speed, no tracking overlay (28_take_video.py). Protocol takes are about 1 min, game takes about 5 min (about 100 s at 3x). The oct game recording holds two takes: game (pulse 1, no trunk compensation) and game_wtc (pulse 2, with trunk compensation). Click a video to play it.

## Slide 4: Three ways to get the trunk angle
Stereo ICP
Stereo depth → YOLO torso mask → chest surface → ICP against the neutral chest
Parameters
3.3 M (seg model) + geometry
Speed
Seg 50 ms on Q6A NPU
Frames covered
Only frames with enough visible chest
SAM 3D Body (teacher)
Meta human-mesh model: one image → full 3D body mesh; trunk = upper-spine joint
Parameters
631 M
Speed
200 ms on RTX 4060 laptop GPU
Frames covered
Every frame
Student (ours)
ResNet18 trained to copy the teacher: grey 256 × 256 crop → keypoints + trunk rotation
Parameters
12.2 M
Speed
4.1 ms on Q6A NPU (W8A16)
Frames covered
Every frame
Notes: Units are kept apart: parameters in millions, speed in milliseconds per frame on named hardware, coverage in percent of frames.

## Slide 5: Definition: total rotation and its error
Start pose: mean trunk orientation over the first 15 frames that camera and mocap both have
  protocol takes: start of the take; game takes: from the GPIO rise
θ(t): angle of the single rotation that takes the trunk from the start pose to its pose at time t
  θ = arccos((trace(R(t) R₀ᵀ) − 1) / 2), 0–180°, never negative
Total rotation error = θ_camera(t) − θ_mocap(t), in degrees; box plots show its absolute value
Needs no camera-to-mocap orientation: a rotation's angle is the same in any axes
Limit: compares sizes only. The right size about the wrong axis scores 0, so the three angles check direction
Notes: Example: Lee (never-seen person), woc take, SAM 3D Body vs mocap. theta is the axis-angle size of the rotation from the start pose: theta = arccos((trace(R(t) R0^T) - 1) / 2). Code: 07_icp_trunk.compare_mocap (mag_i, mag_m).

## Slide 6: Definition: flexion, lateral bend, axial rotation
Trunk axes
Camera: lateral x = left minus right shoulder (3D); anterior a = chest-surface normal; up u = a × x
Mocap: sternum rigid body; origin → lateral marker = x, origin → anterior marker = a, u = a × x
x₀, u₀, a₀: the start pose (mean of the first 15 frames), from the camera geometry
u and x: the same axes after the trunk rotation measured at time t
Mocap in camera axes
Camera and mocap axes differ by a fixed rotation S (marker mounting, calibration)
S is fitted per method by Kabsch: mocap rotation vectors onto the method's rotation vectors, frames > 5° from the start pose
Fitted on the person's woc take, reused for their wc and game takes; game takes move too little to fit S
Total rotation does not use S
Notes: 06_trunk_axis.trunk_angles. Planar projection angles in the neutral trunk axes, not an Euler sequence. The neutral axes come from the stereo geometry and are the same for every method; what differs is the rotation being split: ICP's chest registration, or the c_spine3 joint for SAM 3D Body and the student.

## Slide 7: Definition: how errors are scored
Pairing camera and mocap frames
Sync: the camera's GPIO pin is high while Motive records; its rising edge is mocap t = 0, and each camera frame takes the nearest mocap sample
Mocap is split where the rigid body re-locks (jump > 10 cm); the longest clean stretch is scored
Frames scored: protocol takes every frame (30 Hz); Sep 23 game every 3rd (10 Hz); Oct game every 6th (5 Hz)
Camera angles smoothed (Savitzky–Golay, 11 frames, order 3); mocap left raw
Frames a method did not solve are dropped, never interpolated
Frames solved (%) = frames in the mocap window where the method returned an angle
Statistics
Error per frame = camera − mocap (°); boxes and medians use its size |error|
Box = middle 50 % of frames, line = median, whiskers = 5th–95th percentile
Protocol takes: every method on the frames ICP solved; game takes: each method on its own frames
Frames are pooled across takes; neighbouring frames are correlated, so no significance tests
Mean of per-take medians (student versions): median per take, averaged over the 6 unseen takes
Slope = least-squares line of reading vs mocap; gain (synthetic check) = slope through zero
Seen = in the student's training set; unseen = never trained on · Speed = network time, median of 100 runs
Notes: Sync and teleport split: 07_icp_trunk.compare_mocap. Smoothing: 06_trunk_axis.smooth_series (SAVGOL_WINDOW 11, SAVGOL_POLY 3). Statistics: deck_figs.py in the deck generator.

## Slide 8: ICP is accurate when it sees the chest, not in the game
Box = middle 50 % of frames · line = median (printed) · whiskers = 5th–95th percentile
Flexion / lateral / axial: each method's camera-to-mocap orientation fitted on the person's woc take
Same 7 people, same day (Oct)
Protocol takes: median 1.4–1.7° on every axis
Game takes: median 3.6–4.7°, about 3× worse
In the game the arms reach across the chest and the robot cube sits in front of it
Scored only on frames ICP solved: the frames it lost are not in these boxes
Notes: Same seven people recorded on the same day. Protocol = woc + wc takes. Game = the game take with trunk compensation. ICP is scored only on frames it solved, so the game boxes flatter it.

## Slide 9: In the game, ICP loses up to half the frames
Chart: Frames with a trunk angle, per person (%)
| Arjun | Diwakar | Gokul | Jerome | JS | Lee | Ram
ICP, woc takes | 88 | 96.7 | 99.9 | 92.4 | 93.4 | 98.9 | 96.9
ICP, wc takes | 87.6 | 85.4 | 99.9 | 97.3 | 88.4 | 92.4 | 90.7
ICP, game takes | 72.6 | 49.4 | 67.6 | 65.1 | 88.4 | 47.8 | 88.4
SAM 3D Body, game takes | 100 | 100 | 99.6 | 100 | 99.9 | 99.2 | 100
Protocol takes: ICP solves 85–100 % of frames
Game takes: 48–88 % (Lee 48 %, Diwakar 49 %)
SAM 3D Body: 99–100 % in every take
Lost frames: the visible chest is too small to register, not a stereo failure
  Only 194 of 1312 lost frames (Gokul, Sep) were stereo holes
Notes: Frames solved = in-window frames where ICP returned a pose. SAM 3D Body returns a pose on 99-100 % of frames in every take.

## Slide 10: ICP error is set by how much chest it can see
Chart: Median total-rotation error (°)
| 200–400 | 400–600 | 600–800 | 800–1200 | ≥1200
Stereo ICP | 6.2 | 7.6 | 3.6 | 1.7 | 1.2
SAM 3D Body | 2 | 2.9 | 2.6 | 1 | 1.4
Chart: Share of game frames in each bin (%)
| <200 | 200–400 | 400–600 | 600–800 | 800–1200 | ≥1200
Game frames | 30.9 | 8.4 | 26.9 | 16.5 | 12.1 | 5.3
31 % of game frames have < 200 chest points: ICP cannot run there
≥ 800 points: ICP 1.2–1.7°, as good as the protocol takes, but only 17 % of frames get there
200–800 points: ICP 3.6–7.6°; SAM 3D Body 1.0–2.9° at every size
Notes: Sep 23 game takes, 5 people. Chest points = size of the visible torso surface after removing arms. Both methods scored on the same frames in each bin. ICP cannot run below 200 points.

## Slide 11: Best ICP person in the game: Gokul
Each angle from the start pose · each method vs mocap in its own orientation, set on the woc take · ICP spikes clipped
ICP solves 72 % of frames
Median |error|, flexion / lateral / axial
  ICP 1.5 / 3.0 / 1.3°
  SAM 3D Body 1.5 / 0.9 / 2.1°
ICP's weak axis is lateral bend: frequent ±10° spikes
SAM 3D Body reads 3–5° more axial twist between 50 and 150 s, and misses several flexion peaks
Notes: Whole Sep 23 game take. Each method is compared with mocap split into the three angles using that method's own camera-to-mocap orientation, fitted on the person's woc protocol take and applied to the game take. Fitted on the game take itself (what 21/15 do) it is poorly constrained: the two halves of the take give fits 13-14 degrees apart, and the teacher's game fit for Gokul is 32 degrees from its woc fit.

## Slide 12: Worst ICP person in the game: Kabil
Each angle from the start pose · each method vs mocap in its own orientation, set on the woc take · ICP spikes clipped
ICP solves 50 % of frames
Median |error|, flexion / lateral / axial
  ICP, first 170 s: 3.6 / 27.1 / 5.4°
  ICP, after 170 s: 2.3 / 5.2 / 2.7°
  SAM 3D Body, whole take: 1.7 / 1.5 / 1.6°
ICP fails mostly in lateral bend
Why: next slide
Notes: Whole Sep 23 game take. Each method is compared with mocap split into the three angles using that method's own camera-to-mocap orientation, fitted on the person's woc protocol take and applied to the game take. Fitted on the game take itself (what 21/15 do) it is poorly constrained: the two halves of the take give fits 13-14 degrees apart, and the teacher's game fit for Gokul is 32 degrees from its woc fit.

## Slide 13: Best and worst ICP person in the game: the videos
Gokul best: ICP solves 72 % of frames
Kabil worst: ICP solves 50 % of frames
Kabil: the robot cube sits in front of his chest, so ICP sees only a strip of it. Its median lateral bend error is 27.1° in the first 170 s and 5.2° after
SAM 3D Body sees the whole body: on Kabil its median error stays at or under 2.8° for every angle, before and after 170 s
Both at 3× speed; click a video to play
Notes: Sep 23 game takes, reference camera rotated upright, cut to the mocap window (GPIO pulse), 3x speed (render_sep23_game.py). Same takes as the two trace slides before. Click a video to play it.

## Slide 14: Why Kabil's ICP is wrong: the robot cube covers his chest
Flexion / lateral / axial, degrees from the start pose
Start (t = 1 s): whole chest visible, 1,352 points; every method within 1° of 0
Early play (t = 76 s): the cube covers the chest, a 253-point strip is left
  ICP 15 / -39 / 7 vs mocap 7 / 1 / -7: lateral 40° off
  SAM 3D Body 4 / 2 / -8 vs mocap 7 / 4 / -6
Late play (t = 245 s): cube aside, 685 points
  ICP 8 / 4 / -11 vs mocap 6 / 4 / -8
  SAM 3D Body 10 / 3 / -13 vs mocap 6 / 7 / -7
Notes: Top: undistorted camera frames. Bottom: front view of the chest surface ICP registered (orange) over the neutral chest (grey), camera coordinates in metres. Angles are total rotation from the start pose.

## Slide 15: Fixes tried for ICP: none closed the gap
Fix | Frames solved (%) | Error (°) | Outcome
Mode shell: keep the dominant depth layer | 17–62 → 50–72 | New frames poor: 7–12 median | Kept as default; needs a size gate
Crop to top 20–40 % of the shell | About unchanged | Spikes trimmed (Diwakar p95 27 → 11); protocol worse (2.7 → 4.6) | Dropped
Blend with MediaPipe landmarks | Higher | 13–20 median, r 0.1–0.2 | Dropped: fails on the same frames
Mono depth (YOLO26n-depth) + stereo | 53 → 55; mono only 77 | Mono only: 7.8 median | Dropped
Gate: ≥ 800 chest points only | 1–31 | 0.1–2.4 median, p95 ≤ 5.5 | Accurate but too few frames
Every fix trades frames for accuracy; no geometric variant gets both in the game
The failure is partial-overlap registration of a small chest patch, not depth noise or brightness
Notes: Coverage and accuracy kept in separate columns. Numbers from the Sep 23 game takes unless noted.

## Slide 16: SAM 3D Body: a full body from one image
Explains the arms instead of masking them → an answer on every frame
Trunk = upper-thoracic spine joint (c_spine3), where the sternum marker sits
Input: one camera, grey, undistorted 640 × 400, one fixed person box per take
631 M parameters · 200 ms per frame on an RTX 4060 laptop GPU
Notes: Shown on Lee (never-seen person), woc take. The model explains the arms rather than masking them out, so it answers on every frame.

## Slide 17: In the game, SAM 3D Body beats ICP for 4 of 5 people
Box = middle 50 % of frames · line = median (printed) · whiskers = 5th–95th percentile
SAM 3D Body: median 1.3–3.9° on 100 % of frames
ICP: median 1.4–9.4° on only 50–72 % of frames
ICP wins only for Gokul, who keeps the most chest visible
Teacher limits
  Second camera adds nothing: errors of the two views correlate r 0.96–0.98
  Over-reads large lateral bends by about 20 %
  In the game it reads 1.7–3.9° high for 3 of 5 people; for Diwakar and Gokul more so when the elbows are bent
Notes: Sep 23 game takes. Total rotation needs no axis fit. ICP scored on the frames it solved (50-72 %), SAM 3D Body on all frames.

## Slide 18: The teacher cannot run on the board, so we distil a student
Chart: Network time per frame (ms, log scale)
| SAM 3D Body RTX 4060 GPU | Student ResNet18 Q6A CPU FP32 | Torso seg (ICP) Q6A NPU | Student ResNet18 Q6A NPU W8A16 | MobileNetV2 Q6A NPU W8A8
Time per frame (ms) | 200 | 105 | 50 | 4.1 | 2
15 frames/s budget (66.7 ms) | 66.7 | 66.7 | 66.7 | 66.7 | 66.7
Teacher: 631 M parameters, 200 ms per frame on a laptop GPU; too big for the Q6A
Student labels = teacher output on the training people's recordings + COCO people
Student: 12.2 M parameters (52× smaller)
  4.1 ms on the Q6A NPU, 105 ms on its CPU
Same output on every frame, no depth, no chest visibility needed
Notes: Log scale. Hardware differs per point and is named in the label. The 66.7 ms line is the per-frame budget for 15 frames/s.

## Slide 19: Student problem 1: postures it never saw in training
Box = middle 50 % of frames · line = median (printed) · whiskers = 5th–95th percentile
Flexion / lateral / axial: each method's camera-to-mocap orientation fitted on the person's woc take
v1 trained on 5 people, woc + game takes only
New people, wc takes: lateral 5.0° vs teacher 2.0°
Jerome wc: student 5.0° total vs teacher 1.5°, ICP 1.1°: the error is the student's own
Fix (v2): add wc takes and the Oct sessions of the 4 returning people
Notes: Student v1 was trained on the Sep 23 woc and game takes only. Scored here on the wc takes of the three never-seen people, on frames ICP solved.

## Slide 20: Student problem 2: it learns the people, not the trunk
Box = middle 50 % of frames · line = median (printed) · whiskers = 5th–95th percentile
Flexion / lateral / axial: each method's camera-to-mocap orientation fitted on the person's woc take
People it trained on: student ≈ teacher (flexion 1.4 vs 1.5°, axial 1.3 vs 1.5°)
New people: v2 falls behind (flexion 2.5 vs 1.5°, axial 2.5 vs 2.0°)
Same network, same task: it learned the 5 people, not trunk pose
With synthetic people (v3), new-people flexion 2.5 → 1.9°
v3 is still 0.3–0.4° behind the teacher on new people
Notes: Seen = the four people in the training set (their Oct protocol takes were training data). Unseen = the three test people. Teacher shown on the same frames for reference.

## Slide 21: Student problem 3: it shrinks large flexion on new people
Chart: Reading vs true flexion (°)
| -17.5 | -12.5 | -7.5 | -2.5 | +2.5 | +7.5 | +12.5 | +17.5 | +22.5 | +27.5
Perfect | -17.5 | -12.5 | -7.5 | -2.5 | 2.5 | 7.5 | 12.5 | 17.5 | 22.5 | 27.5
Stereo ICP | -14.1 | -9.7 | -5.8 | -1.7 | 1.7 | 6.7 | 12.4 | 17.3 | 20.3 | 27.8
SAM 3D Body | -15.2 | -11.5 | -5.6 | -1 | 1.9 | 6 | 11.2 | 16.4 | 21 | 24.8
Student v2 (real only) | -13.9 | -10.1 | -5 | -1.1 | 0.5 | 3.7 | 7.6 | 12.8 | 17.9 | 23.9
Student v3 (+ synthetic) | -15.9 | -11 | -5.4 | -1.5 | 0.7 | 4.7 | 9 | 14 | 19.6 | 24.8
Teacher and ICP follow the diagonal (slope 0.92, 0.91)
Student v2: slope 0.81; at 12.5° true flexion it reads 7.6°
Flexion = leaning toward the camera, the weakest cue in one image
Training poses are mostly upright (71–85 % within ±5°), so an unsure student falls back to upright
v3 (+ synthetic): slope 0.85, reads 9.0°: partly fixed
Notes: Unseen people, protocol takes, frames ICP solved. Frames grouped in 5 degree bins of mocap flexion; each line is the median reading in the bin. The dashed grey line is a perfect reading.

## Slide 22: Why render synthetic humans
The bottleneck is identities: only 5 real training people
1,500 random adults, 15,000 frames, 1.6 h on the laptop GPU
Real seated poses: teacher fits of the training people only (52k-pose pool)
Varied rooms, tables, lighting, robot cube, bystanders; camera placed like the rig
Converted to the camera's grey look before labelling
Weak spot: about 10 outfits, so clothing variety is low
Notes: Blender + MPFB2 (MakeHuman, CC0 assets). Poses come only from the training people, so no test-person pose leaks into training.

## Slide 23: Teacher misreads synthetic flexion, so labels use rendered truth
Teacher on synthetic people: flexion gain 0.54 (r 0.63); lateral 0.92, axial 0.98
On real people the teacher reads flexion at gain ≈ 0.9 vs mocap, so this is a synthetic-domain gap; copying it would teach the student to under-read flexion
So: synthetic trunk label = rendered truth (offset to the teacher's mean per person); keypoints and arms stay teacher labels
Notes: Each panel: rendered truth vs teacher reading, relative to each person's mean pose. Gain = least-squares slope through zero.

## Slide 24: Synthetic people improve flexion on new people
Chart: Flexion (°)
| v2 real only | v3 + synthetic | v3 synth 40 % | v3 rerun | v4 flex-balanced
Student | 2.5829 | 2.2114 | 1.9429 | 1.9367 | 1.9643
SAM 3D Body | 1.5588 | 1.5588 | 1.5588 | 1.5588 | 1.5588
Stereo ICP | 1.5359 | 1.5359 | 1.5359 | 1.5359 | 1.5359
Chart: Lateral bend (°)
| v2 real only | v3 + synthetic | v3 synth 40 % | v3 rerun | v4 flex-balanced
Student | 2.253 | 2.7474 | 2.2562 | 2.2808 | 2.4355
SAM 3D Body | 2.333 | 2.333 | 2.333 | 2.333 | 2.333
Stereo ICP | 2.6594 | 2.6594 | 2.6594 | 2.6594 | 2.6594
Chart: Axial rotation (°)
| v2 real only | v3 + synthetic | v3 synth 40 % | v3 rerun | v4 flex-balanced
Student | 2.839 | 2.3451 | 2.3685 | 2.2801 | 2.421
SAM 3D Body | 2.1524 | 2.1524 | 2.1524 | 2.1524 | 2.1524
Stereo ICP | 2.0057 | 2.0057 | 2.0057 | 2.0057 | 2.0057
Chart: Total rotation (°)
| v2 real only | v3 + synthetic | v3 synth 40 % | v3 rerun | v4 flex-balanced
Student | 2.8485 | 2.7637 | 2.6655 | 2.486 | 2.6565
SAM 3D Body | 2.1388 | 2.1388 | 2.1388 | 2.1388 | 2.1388
Stereo ICP | 1.8941 | 1.8941 | 1.8941 | 1.8941 | 1.8941
Adding synthetic people: flexion 2.58 → 2.21°
Synthetic at 40 % of batches: flexion 2.21 → 1.94°, total 2.85 → 2.67°
Same recipe trained twice differs by up to 0.2°
v4 flexion-balanced synthetic poses: flexion 1.96°, no gain
Lateral: worse in v3 (2.75°), back to v2's level with 40 % synthetic (2.26°)
Teacher (dashed blue) is still the ceiling
Notes: Mean over the 6 unseen protocol takes of each take's median error vs mocap, frames ICP solved. 'Rerun' is the same recipe trained again. wc takes split with each method's woc-take orientation.

## Slide 25: Protocol take close-up: Lee (new person)
Each column: one method vs mocap in that method's own orientation · each angle from the start pose · student never saw this person
20 s of the 56 s take (7–27 s)
Median |error| here, flexion / lateral / axial
  ICP 1.1 / 2.6 / 1.0°
  SAM 3D Body 2.6 / 2.6 / 1.1°
  Student v3 2.2 / 2.3 / 3.1°
ICP solves 98 % here and follows every peak
Forward lean at 21 s (mocap 28°): ICP 24°, SAM 3D Body 20°, student 18°
After a 0.3 s mocap dropout at 16.5 s, all three methods' lateral error grows by 2–3° at once: most likely the mocap re-locked, not the cameras
Notes: Lee's woc take, 7–27 s of 56 s, picked as the 20 s where all three angles move most. Student v3 = r18_v3mix_all, which never saw Lee. Each method is compared with mocap in its own camera-to-mocap orientation, fitted on this take. Whole-take median |error| flexion / lateral / axial: ICP 1.0 / 2.9 / 0.9°, SAM 3D Body 1.8 / 2.0 / 1.1°, student 2.2 / 1.6 / 2.2°; ICP solves 99 % of the take.

## Slide 26: Protocol take close-up: Arjun (new person)
Each column: one method vs mocap in that method's own orientation · each angle from the start pose · student never saw this person
20 s of the 51 s take (10–30 s)
Median |error| here, flexion / lateral / axial
  ICP 0.9 / 2.0 / 1.3°
  SAM 3D Body 1.2 / 2.5 / 3.1°
  Student v3 1.3 / 2.8 / 3.7°
ICP solves 81 % here: it loses the ends of both left twists (−57°, −53°) and reads 22° at the +53° twist
Side bend at 10.5 s (mocap 28°): SAM 3D Body 20°, student 23°
SAM 3D Body and student over-read the left twists by 5⁠–⁠7°
Notes: Arjun's woc take, 10–30 s of 51 s, picked as the 20 s where all three angles move most. Student v3 = r18_v3mix_all, which never saw Arjun. Each method is compared with mocap in its own camera-to-mocap orientation, fitted on this take. Whole-take median |error| flexion / lateral / axial: ICP 0.8 / 1.5 / 0.8°, SAM 3D Body 1.2 / 2.5 / 2.5°, student 1.1 / 2.5 / 2.4°; ICP solves 88 % of the take.

## Slide 27: Protocol take close-up: Jerome (new person)
Each column: one method vs mocap in that method's own orientation · each angle from the start pose · student never saw this person
20 s of the 54 s take (17–37 s)
Median |error| here, flexion / lateral / axial
  ICP 1.5 / 1.9 / 1.9°
  SAM 3D Body 1.6 / 0.5 / 1.6°
  Student v3 1.3 / 1.2 / 1.9°
At full twist (54°, 21 s) ICP reads 25° and jumps to about 30° flexion for 2 s while mocap reads 4°: a wrong registration
SAM 3D Body follows lateral bend within 0.6°
Student under-reads the 38° lean at 36 s (32°) and the 41° twist at 23 s (31°)
Notes: Jerome's woc take, 17–37 s of 54 s, picked as the 20 s where all three angles move most. Student v3 = r18_v3mix_all, which never saw Jerome. Each method is compared with mocap in its own camera-to-mocap orientation, fitted on this take. Whole-take median |error| flexion / lateral / axial: ICP 1.2 / 1.4 / 1.4°, SAM 3D Body 1.5 / 0.5 / 1.5°, student 1.7 / 1.1 / 1.2°; ICP solves 92 % of the take.

## Slide 28: Where we are: new people, protocol takes
Box = middle 50 % of frames · line = median (printed) · whiskers = 5th–95th percentile
Flexion / lateral / axial: each method's camera-to-mocap orientation fitted on the person's woc take
ICP is the most accurate where it sees the chest: 1.4–2.2°
Teacher 1.5–2.0°
Student v3 1.9–2.7°, student v2 2.0–2.8°
Student v3 vs teacher: +0.3° flexion, +0.3° axial, +0.8° total
Synthetic people closed most of the flexion gap (v2 → v3)
Notes: Never-seen people (Arjun, Jerome, Lee), woc + wc takes, frames ICP solved (about 90 % of frames).

## Slide 29: Where we are: new people, game takes
Box = middle 50 % of frames · line = median (printed) · whiskers = 5th–95th percentile
Flexion / lateral / axial: each method's camera-to-mocap orientation fitted on the person's woc take
ICP solves only 48–73 % of frames; teacher and student 99–100 %
ICP: 3.7–5.1° median even on the frames it solved
Teacher 1.9–2.9° · student v3 2.1–3.3°
Synthetic people help axial most here: 3.9 → 3.3°
Student v3 minus teacher: flexion −0.2, lateral +0.2, axial +0.6, total +0.2°
Notes: Never-seen people, game take with trunk compensation, every frame each method solved. ICP boxes leave out the frames it lost.

## Slide 30: Summary
Stereo ICP
Best accuracy when the chest is visible (1.4–2.2° on new people's protocol takes). In the game it loses up to half the frames, and small-shell frames are silently wrong.
SAM 3D Body
Answers on every frame, 1.3–3.9° in the game. 631 M parameters, 200 ms on a laptop GPU: cannot run on the Q6A.
Student
4.1 ms on the Q6A NPU. v1 failed on unseen postures; real-only v2 overfit its 5 people and shrank large flexion.
Synthetic humans
1,500 new identities with exact trunk truth. New-people flexion 2.58 → 1.94°; still 0.3–0.8° behind the teacher.
Options from here
  Use ICP when the chest shell is large (≥ 800 points), the student otherwise
  More identity variety: outfits, body shapes
  Mocap as labels for the training people (beat the teacher)
  Per-axis gain correction for flexion
Compare recipes with ≥ 2 runs each: one run cannot resolve 0.2°